## Samples
bar_name | sample_id | sample_name | project_name | institution | proposal_id | bar_spot | front | grazing | angle | height | sample_priority | notes | location | bar_loc | acq_history
20250923_Test | AF1_front | AF1_front | calibration | NIST | 318915 | 0A | TRUE | FALSE | 0 | 0 | 1 | Top fiducial hole used to calibrate sample bar position. Don't change this row.
20250923_Test | rsoxs_sample | rsoxs_sample | Ketkar | NIST | 318915 | 1A | TRUE | FALSE | 0 | 0.25 | 1 | Typical transmission RSoXS sample on SiN membrane that is mounted to the front of the bar. 0 degrees angle corresponds to normal incidence. bar_spot can be filled in after you mount samples.
20250923_Test | TEY_NEXAFS_sample | rsoxs_sample | Ketkar | NIST | 318915 | 20A | TRUE | TRUE | 90 | 0.5 | 1 | Typical TEY NEXAFS sample on silicon wafer. For grazing samples, 90 degrees corresponds to normal incidence.
20250923_Test | SiN_Blank | SiN_Blank | calibration | NIST | 318915 | 18A | TRUE | FALSE | 0 | 0.25 | 1 | Blank SiN membrane to calibrate scans for other samples mounted onto SiN. Make sure the thickness and other characteristics are identical to the membranes used for other samples.
20250923_Test | SBA15 | SBA15 | calibration | NIST | 318915 | 28A | TRUE | FALSE | 0 | 0.25 | 1 | For SDD and beam center calibration. Don't change this row.
20250923_Test | HOPG | HOPG | calibration | NIST | 318915 | 30A | TRUE | TRUE | 90 | 0.5 | 1 | Energy calibrant. Don't change this row.
20250923_Test | AF2_front | AF2_front | Calibration | NIST | 318915 | 0A | TRUE | FALSE | 0 | 0 | 1 | Bottom fiducial hole used to calibrate sample bar position. Don't change this row.
20250923_Test | OpenBeam | OpenBeam | calibration | NIST | 318915 | 40A | TRUE | FALSE | 0 | 0 | 1 | Bar moved out of beam path for reference scan. Don't change this row.

## Acquisitions
sample_id | configuration_instrument | scan_type | energy_list_parameters | polarization_frame | polarizations | exposure_time | exposures_per_energy | cycles | sample_angles | spiral_dimensions | group_name | priority | notes | acquire_status | uid_local
OpenBeam | NoBeam | time2D |  |  |  |  | 3 |  |  |  | Commissioning | 1 | Delete this row for beamline users. For beamline scientist commissioning purposes only. Use duplicate_sample to duplicate OpenBeam such that the sample manipulator motor position is that when the sample is all the way retracted. Replace sample_id here. May need to add a fake bar_loc column with x0 and y0 replaced with x and y positions for rotation to work. Run a counting scan (no motors are moved, and only recording data over time) to test beamline workflows.
HOPG | WAXSNEXAFS | nexafs | carbon_NEXAFS | lab | [90] | 1 |  | 1 | [20] |  | BeamCalibration | 2 | Do not delete this row. TEY NEXAFS (scan_type = nexafs) on HOPG (sample_id = HOPG). 2D detector is retracted, but beamstop is in position (configuration_instrument = WAXSNEXAFS) --> Protects 2D detector from sample bar rotation and reflected X-rays. The sample is rotated to an angle of 20 (angle between incident beam and sample is 20 degrees), sample_angles = [20] --> 20° angleensures a larger beam footprint on HOPG. A carbon-edge energy scan is performed at 90° polarization --> The beam energy was calibrated at 90° polarization, so the calibration is being validated at this energy. The energy scan is performed in ascending and descending energy directions (cycles = 1) --> This checks for reproducibility of the NEXAFS profiles. Optional: Run at additional polarizations (e.g., polarizations = [0, 20, 30, 55, 70, 90]) and/or sample angles (sample_angles = [20, 30, 55, 70, 90]) to see polarization dependence of HOPG.
SBA15 | WAXS | rsoxs | (200, 91.65, 291.65, 8.35, 300, 100, 1300) | lab | [90] | 0.1 |  |  |  |  | BeamCalibration | 3 | Do not delete this row. RSoXS energy scan (scan_type = rsoxs) on SBA15 (sample_id = SBA15). 2D detector and beamstop are in the beam path (configuration_instrument = WAXS) --> Scattering images can be captured. Scan is over a custom list of energies. Between 200 and 291.65 eV, scans will be taken in 91.65 eV steps. Between 291.65 and 300 eV, scans will be taken in 8.35 eV steps. Between 300 and 1300 eV, scans will be taken in 100 eV steps. 0.1 s exposure times per image (exposure_time = 0.1) --> Reduces saturation of 2D detector. The scan is performed at 90° polarization (polarizations = [90]), cycles = 0 (default, only runs in ascending energy order but no descending energy order), and sample_angles = [0] (beam is at normal incidence from sample surface for transmission sample). Optional: Use multiple exposures (e.g., exposures_per_energy = 50) to increase signal-to-noise to assess more subtle features in the sample and/or 2D detector artifacts.
rsoxs_sample | WAXS | spiral | 270 | lab | [0] | 1 |  |  |  | [0.3, 1.8, 1.8] | Spirals | 4 | Rows from here and below can be modified, duplicated, or deleted as needed unless specifically indicated otherwise. Typical spiral scan at 270 eV to survey different spots on your sample. The spiral dimensions correspond to 0.3 mm step size between spots, 1.8 mm width of the full spiral scan, and 1.8 mm height of the full spiral scan. In most cases, width and height should be the same.
rsoxs_sample | WAXSNEXAFS | nexafs | carbon_NEXAFS | lab | [0] | 1 |  | 5 |  |  | TransmissionNEXAFS | 5 | Typical transmission NEXAFS scan across the carbon edge. Cycles = 5 indicates that it will scan up and down in energy 5 times.
rsoxs_sample | WAXSNEXAFS | nexafs | oxygen_NEXAFS | lab | [0] | 1 |  |  |  |  | TransmissionNEXAFS | 6 | Typical transmission NEXAFS scan across the oxygen edge.
rsoxs_sample | WAXSNEXAFS | nexafs | nitrogen_NEXAFS | lab | [0] | 1 |  |  |  |  | TransmissionNEXAFS | 7 | Typical transmission NEXAFS scan across the nitrogen edge.
TEY_NEXAFS_sample | WAXSNEXAFS | nexafs | carbon_NEXAFS | lab | [0] | 1 |  |  | [20, 30, 55, 70, 90] |  | TEY_NEXAFS | 8 | TEY NEXAFS where a series of carbon edge scans are taken at different sample angles
TEY_NEXAFS_sample | WAXSNEXAFS | nexafs | carbon_NEXAFS | lab | [0, 20, 30, 55, 70, 90] | 1 |  |  | [20] |  | TEY_NEXAFS | 9 | TEY NEXAFS where the sample is rotated to an angle of 20 (angle between incident beam and sample is 20 degrees) and a series of carbon edge scans are taken at different beam polarizations.
rsoxs_sample | WAXS | rsoxs | carbon_RSoXS | lab | [0, 90] | 0.1 |  |  |  |  | RSoXS | 10 | Typical RSoXS scan across carbon edge at 0 and 90 degrees polarization. Exposure time per energy is 0.1 s.
rsoxs_sample | WAXS | rsoxs | oxygen_RSoXS | lab | [0, 90] | 1 |  |  |  |  | RSoXS | 11 | Typical RSoXS scan across oxygen edge at 0 and 90 degrees polarization
rsoxs_sample | WAXS | rsoxs | nitrogen_RSoXS | lab | [0, 90] | 1 |  |  |  |  | RSoXS | 12 | Typical RSoXS scan across nitro edge at 0 and 90 degrees polarization
rsoxs_sample | WAXS | rsoxs | (270, 1, 280, 0.1, 300, 0.5, 320) | lab | [0, 90] | 1 |  |  |  |  | RSoXS | 13 | RSoXS scan for a custom list of energies. Between 270 and 280 eV, scans will be taken in 1 eV steps. Between 280 and 300 eV, scans will be taken in 0.1 eV steps. Between 300 and 320 eV, scans will be taken in 0.5 eV steps. The list should contain at least 3 items, but no maximum limit.
rsoxs_sample | WAXS | rsoxs | 285 | lab | [0, 90] | 1 | 200 |  |  |  | RSoXS | 14 | RSoXS scan at which 200 images are taken at 285 eV, both at 0 and 90 degrees polarization.